GMW

USER’S MANUAL

MODEL:  5203

PROJECTED VERTICAL FIELD MAGNET

Date Sold:  _______________

Serial number:  ___________

PROPRIETARY

THIS DOCUMENT CONTAINS CONFIDENTIAL INFORMATION PROPRIETARY TO GMW ASSOCIATES.  IT MUST NOT BE REPRODUCED OR DISCLOSED TO OTHERS OR USED IN ANY WAY EXCECPT FOR THE INSTALLATION, OPERATION OR MAINTENANCE OF GMW ASSOCIATES PRODUCTS.

File No: M5203_Proj_Vert_Field_Magnet_Feb_17_2017.pdf		  	   Revision Date: February, 2017

| GMW | 955 Industrial Road, San Carlos, CA 94070 Tel: (650) 802-8292 Fax: (650) 802-8298 Email: tom@gmw.com Web site: http://www.gmw.com |
| --- | --- |

- CONTENTS

1.	SAFETY	3

1.1.	Symbols Used In This Manual	3

1.2.	Installation, Operation and Service Precautions	4

1.3	Lockout/Tagout Procedures	5

2.	SPECIFICATIONS	8

2.1.	General Specifications	8

2.2.	Model 5203 Electrical and Water Connections	9

2.3.	Drawing 11907-0224-0_S2_F_PVFM_General_Assembly	10

2.4.	Drawing 11907-0233-0_E_5203_PVFM_Terminal_Block_Assembly	11

3	WARNINGS	12

4	INSTALLATION	14

4.1	Unpacking Instructions and Damage Inspection	14

4.2	Siting Considerations	15

4.3	Electrical Circuit	16

4.4	The 5203 Interlocks	17

4.5	Cooling	18

4.6	Mounting Plate	19

5	OPERATION	20

5.1	General	20

5.2	Field Control Operation	21

6	MAINTENANCE	22

6.1	Cooling Circuit	22

6.2	Strain Relief Check	22

6.3	Preventative Maintenance	22

7	STANDARD OPTIONS	23

7.1	Standard Pole Options	23

8	CUSTOM OPTIONS	24

9	EXCITATION CURVES AND FIELD UNIFORMITY	25

10	TEST DATA	40

11	WIRING DRAWINGS	41

11.1	Drawing 13907-0025-1_B 5203 1xBOP20-50 and PDU System Wiring	41

11.2	Drawing 13907-0026-1_B 5203 2xBOP20-50 and PDU System Wiring	42

11.3	Drawing 16907-0116-1-S1_A Cable Current Interlock 5203 and PDU to  BOP20-50	43

12	APPENDIX	44

12.1	Recommended Mains Connection	44

# SAFETY

## Symbols Used In This Manual

|  | Important information. This label indicates important information that must be read and understood prior to operation. |
| --- | --- |
|  | Warning relating to the presence of magnetic field. |
|  | Warning relating to a trip hazard. |
|  | Warning relating to hot surfaces. |
|  | Warning relating to the potential for electrical shock. |
|  | Warning relating to the potential of injury due to lifting heavy weight. |

## Installation, Operation and Service Precautions

|  | During installation or servicing this magnet must be locked out following the Lockout/Tagout procedure defined in section 1.3. |
| --- | --- |
|  | This product is designed for use in accordance with SOP39. Hazardous voltages are present within this product during normal operation. The product should never be operated with any of the covers removed. |
|  | When operated at peak direct current of 63A the surface temperature of the coils of the electromagnet can rise as high as 80°C. Care must be taken not to touch the coils for extended periods. |

## Lockout/Tagout Procedures

Purpose:

To protect personnel and ensure that machines and equipment are isolated from potentially hazardous electrical energy. Lockout or tagout must occur before employees perform service, maintenance, or renovation. This is important where unexpected start-up could cause personal injury, fire, or equipment damage.

Policy:

All equipment shall be locked out where possible.  Where such control is not possible, equipment may be tagged out-of-service.  In all instances, equipment shall be made inoperable to protect against possible operation where such operation may cause personal injury or damage. Employees must not attempt to operate any switch or source of energy which is locked out or tagged out.

Procedures:

- When working on systems which could accidentally be activated, the system shall be locked out or tagged out by use of a safety lockout device and padlock. In addition, a tag shall be used to identify the purpose of the shutdown, the employee involved, the date the unit was removed from service, and when the system may operate again.

- If more than one source of energy is present, all such sources must be locked out or tagged out. Special procedures must be followed to ensure that the equipment is disconnected from an energy management system or emergency generator system that may start or energize the equipment.

- A lockout is required on all systems where possible. A tagout is an acceptable means of protection on systems which are less hazardous. An example of less hazardous is a device, if started accidentally, would not cause personal injury.

- If more than one person is involved in the repair, each person shall install a lock and or tag to the equipment energy source. An employee may not use the tag or lock of another employee.

- After servicing, renovation, or maintenance is complete, the area must be checked for tools, parts, removed guards, and assurance that no personnel are in the danger zone. Then the lockout or tagout will be removed by the same employee who initially locked it out so energy may be restored to the equipment.

- If the employee who locked out or tagged out the equipment is unavailable, the supervisor may remove the lock or tag if the following conditions are met:

1.   Verification that the employee who applied the device has left for the duration of the shift and is not at the job site.

2.   Made reasonable efforts to reach the employee.

3.   Inform the employee that the lock or tag has been removed and the system is no longer de-energized, before the employee resumes work.

- Employee's using lockout/tagout devices shall have training about this program, and shall have annual retraining to ensure that the employee understands and follows this program.  The training and retraining shall be documented with the training records maintained by the training coordinator.

- Outside contractors are required to follow this policy or provide a similar policy that is in compliance with Occupational Safety and Health Administration (OSHA) Standard 1910.147.  Under no circumstances are outside contractors authorized to remove a lockout/tagout device nor are they allowed to energize a locked out/tagged out system.

Lockout/Tagout Procedures Checklist:

The following steps must be followed in sequence to properly lockout/tagout and re-establish energy:

I. Understand the hazard:

Electrical

Electrical Shock and or burn could result from contact with the exposed conductors line voltage or high voltage equipment. Flying parts or fire could result if this circuit were shorted.  Electricity should be controlled at the circuit breaker, main switch, or fuse box.

Mechanical

Equipment or machinery can inflict tissue or skeletal injury through crushing, laceration or impalement. This can be controlled through the main electrical switch, plug, circuit breaker or anti-motion pin.

Thermal

Can cause burns or fires. It can be controlled by the main electrical switch, electrical plug control, electrical circuit breaker, electrical fuse box, steam valve, fluid line valve or shielding.

II. Shutdown:

Know what type of energy the machine uses.

Identify its potential hazards.

Find the switches or other devices that control energy and need to be locked out.

Let employees know that you will be locking or tagging out the equipment and why.

Turn off the machine or equipment.

III. Isolate the source of energy

Electrical

- Locate the main switch box or circuit breaker.

- Open the breaker, open the switch or remove the plug.

- Attach a lockout enabling device if the circuit cannot otherwise accommodate a padlock.

- Place plug in a plug lock box.

- Note, the center of the grip of the operating handle of the disconnecting means, when in its highest position, shall be not more than 2.0 m (6 ft 7 in.) above the servicing level, nor should they be less than 0.6m (2 ft) above the servicing level and be within easy reach of the normal working position of the operator. It should be readily accessible with doors in the open or closed position and be located within line of sight from the operator station.

- The disconnect should be readily identifiable as an isolating means and marked to identify the equipment that it is disconnected.

Mechanical/Storage Potential Energy

- Lockout enabling device.

- Secure the energy controlling lockout by attaching a personal lock and completed tag to the lockout enabling device. If more than one person will be performing the work, each must apply his own lock to a multiple lock device.

- Release all stored energy.

- If there is a heat exposure, allow to cool.

Release from Lockout/Tagout

1.  Inspect the surrounding area following completion of work for loose tools, parts, correct valve settings, system integrity, exposed conductors.

2.  Check that all machine guards are in place and reconnected if applicable.

3.  Notify others in the area that the equipment is about to be made operational and returned to service.

4.  Remove personal lock, tag, and lockout enabling device. This step must be performed by the same person who applied the tag and lock.

Lockout/Tagout Training

The lockout/tagout program is designed to train employees on disabling powered equipment from their power sources before beginning any servicing or maintenance work. Lockout/tagout training is required for all employees who may possibly need to lockout and tagout equipment.

Employees will receive annual training in the following areas:

1.  Recognition of applicable hazardous energy sources.

2.  Methods and Means necessary for energy isolation and control.

3.  Restrictions and limitations of lockouts.

# SPECIFICATIONS

## General Specifications

| Projected Field (X, Y, Z= 0, 0, 5mm) at: Max triangle wave current 100A, 945W Max sinusoidal current 85A, 945W Max DC current 63A, 945W | Bz = ± 0.518T (5180G) Bz = ± 0.476T (4760G) Bz = ± 0.407T (4070G) |
| --- | --- |
| Coils (series connection) |  |
| Coil Resistance (20°C) | 0.173 Ohm |
| Max Resistance (80°C) | 0.240 Ohm |
| Max Power (water cooled @ 8 lpm) | 63A/15V (945W) |
| Self Inductance (low field) Self Inductance (high field) (The apparent inductance increases with frequency due to eddy currents in the solid poles) | 3500µH at 1 Hz 895µH at 1 Hz |
| Water Cooling (18°C) | 4.0 bar (60 psid), 8 liters/min (2.1 US GPM) |
| Over-temperature Interlock | Selco UP62-080C thermostat Contact rating 250VAC, 2.5A. Closed below 80°C |
| Water flow interlock | GEMS FS-380. Closed above 1.5 US GPM |
| Dimensions | Drawing 11907-0224-0_F 74.0 mm W x 74.0 mm D x 123.5 mm H (2.91 inch W x 2.91 inch D x 4.86 inch H) |
| Mass | 2.5 kg (5.5 lb) |

|  | CAUTION - The value of maximum coil resistance given should not be exceeded. At this resistance the coils are at maximum safe temperature for continuous operation. |
| --- | --- |

## Model 5203 Electrical and Water Connections

| DC Current | The 5203 electromagnet is shipped with the power and water connections between the junction box and the magnet head connected. In the event that these connections are broken, the correct reconnections can be made by referencing drawings 11907-0224-0_S2_F and 11907-0233-0_S1_E. |
| --- | --- |
| Ground | The coil cooling plates, the magnet body and the base plate are all earthed through the power supply earth. |
| Interlocks (refer Section 8) | The Model 5203 has a single thermostat, Selco part no UP62-080C. It is located between the base yoke plate and the lower cooling plate. The thermostat is normally closed, opening when the coil heatsink temperature exceeds 80°C, ± 5°C. In addition, a water flow switch is located in the junction box. This is normally open and will close once the flow rate rises above 1.5 US GPM. |
| Water | Clean, cool (16°C - 20°C) water at 8.0 lpm at 4.0 bar (60 psid) should be used to cool the 5203 magnet. This can be provided by house water supply or from a recirculating chiller. An appropriate recirculating chiller is the TF90 from Thermofisher. The magnet is shipped with a cable junction box connected, water connections should be made to the two connections on this junction box. The flow direction is indicated on the side of the flow switch. Refer to drawing 11907-0224-0_S2_F. If house water supply is used, an in-line 5 micron water filter should be used for the supply line to the magnet. Also, 10m of hose is provided to give two 5m lengths for connection between the magnet and the recirculating chiller (if ordered), brass fittings on the chiller are already provided. Note that 4 hose clamps have also been provided for fitting these water hoses. The power cable that connects the power supply to the magnet (16907-0116-1-S1_A) is also provided. Outlet: 1/4 inch Swagelok Inlet: 1/4 inch Swagelok |
| CAUTION | Ensure that the high current connections are tight. Loose connections may lead to oxidation and overheating. The field stability may be degraded and the current terminations damaged. |

## Drawing 11907-0224-0_S2_F_PVFM_General_Assembly

## Drawing 11907-0233-0_E_5203_PVFM_Terminal_Block_Assembly

# WARNINGS

REFER TO WARNINGS BELOW BEFORE OPERATING ELECTROMAGNET

| 1 Personnel Safety |  |
| --- | --- |
|  | The Model 5203 has a low level of fringe field. In operation the magnet fringing field can be in excess of 0.5mT (5G). This can cause malfunctioning of heart pacemakers and other medical implants. We recommend that the fringing field should be mapped and warning signs be placed outside the 0.5mT (5G) contour. Entry to this region of higher field should be restricted to qualified personnel. |
| 2 Ferromagnetic Objects |  |
|  | During operation the magnet exerts strong magnetic attraction towards ferromagnetic objects in the near vicinity. Loose objects can be accelerated to sufficient velocity to cause severe personnel injury or damage to the coils. Keep ferromagnetic tools clear. |
| 3 Arcing |  |
|  | This magnet stores considerable energy in its field during operation. Do not disconnect any current lead while under load or the magnetic field energy will be discharged across the interruption causing hazardous arcing. |
| 4 Coil Hot Resistance |  |
|  | Do not exceed the maximum coil hot resistance given in the specifications or coil overheating and possible damage may occur. |
| 5 Interlocks |  |
|  | These should always be connected if the magnet is to be operated unattended, to avoid the possibility of coil overheating caused by excessive power dissipation or inadequate cooling. |

| 6 Watches, Credit Cards, and Magnetic Disks |  |
| --- | --- |
|  | Do not move magnetically sensitive items into the close vicinity of the magnet. Even some anti-magnetic watches can be damaged when placed in close proximity to the magnet during operation. Credit cards, and magnetic disks are affected by magnetic fields as low as 0.5mT (5G). |
| 7 Trip Hazards |  |
|  | Caution must be taken when moving around the magnet as there are potentially cables on the floor that will present a trip hazard. It is highly advised that floor conduit be purchased that will cover the cables. This will also protect the cables from potential damage. |

# INSTALLATION

|  | Caution: Take care when handling system. All movement, lifting and installation of the 5203 Electromagnet must be under the supervision of an experienced person to prevent the possibility of serious injury or damage to the electromagnet and associated equipment. |
| --- | --- |

## Unpacking Instructions and Damage Inspection

To unpack the electromagnet please use the following procedure.

1.	First remove the wire clips and screws from the ‘top crate cover’ of the shipping crate.

2.	Remove the top crate cover and set it aside. Remove the protective foam packaging material.

- Inspect the magnet and junction box to ensure that no damage has occurred during shipment.  If damage is evident report the damage in detail to the shipper for claim and simultaneously notify GMW in case assessment of the damage must be made.  If no damage is found proceed with magnet unpacking and installation.

- Carefully lift magnet, cables, hoses and junction box clear of the shipping crate.

- All conductors of the same ac circuit routed to the same location shall be contained in the same raceway. Conductors external to the electrical equipment enclosure(s) shall be enclosed in raceways (ducts).

## Siting Considerations

|  | The Model 5203 is not magnetically shielded. Ferro-magnetic material in the vicinity of the magnet will modify the magnitude and uniformity of the central region magnetic field. As a general rule avoid magnetic material closer than approximately 1 meter of the central region. |
| --- | --- |

Background fields such as the geomagnetic field and alternating field from 60Hz power sources are not shielded by the magnet and will add to the field produced by the magnet.  If possible these background fields should be measured and their effects evaluated before the Model 5203 magnet is installed.  It may be necessary to orient the Model 5203 axis to minimize the effects of external fields, to relocate ac power sources or to install suitable magnetic shielding.

The 5203 Electromagnet features four M3 clearance holes in its base plate and should be securely mounted to a stable foundation. Refer Drawing 17907-0412-0_C in section 4.6

The width of the working space in front of control cabinets and compartments should be the width of the opening or 762mm (30 in.), whichever is greater. Where control equipment or devices are mounted on or through the fixed area around the opening into the control cabinet or compartment, the width of the working space in front of the control cabinet or compartment shall include the width of the fixed area containing the control equipment and devices.

The working space height shall be clear and extend from the grade, floor, or platform to a height of 2.0 m (6 ½ ft). Within the height requirements of Section 11.5, other equipment associated with the machine located above or below the control cabinet or compartment shall be permitted to extend not more than 150 mm (6 in.) beyond the front of the electrical control cabinet or compartment.

The working space shall permit at least 90-degree opening of control cabinet and compartment doors or hinged panels.

The main disconnect plug socket combination shall be marked “Main Disconnect. The Main Disconnect should be located within sight from the operator station and be readily accessible. The height of Main Disconnect must not be greater than 2m height from floor.

The switch rated plug/receptacle shall not be placed below 0.6m to be compliant with NFPA 79.

At least one entrance of sufficient area shall be provided to give access to the working space around control cabinets or compartments and working space should not be used for storage.

## Electrical Circuit

|  | Never connect or remove cables from the magnet with the power supply connected. The stored energy in the magnet can cause arcing resulting in severe injury to personnel or equipment damage. |
| --- | --- |

The magnet is shipped with a power and interlock cable attached. The recommended current cable for the 5203 is shown in Drawing 16907-0116-1-S1_A (section 12). Because the magnet stores a significant amount of energy in its magnetic field, special care should be taken to ensure that the current terminations are secure and cannot work loose in operation.  Local heating at the terminations can cause rapid oxidation leading to a high contact resistance and high power dissipation at the terminals.  If left unattended this can cause enough local heating to damage the terminals and the coils.

## The 5203 Interlocks

The Model 5203 has one Selco UP62-080C thermostat. It is located between the base yoke plate and the lower cooling plate. The thermostat is normally closed, opening when the coil heatsink temperature exceeds 80°C, ± 5°C.

The interlock circuit also includes a flow switch (GEMS FS-380) located at the junction box. This flow switch is open for flow rates above 5.4 lpm (1.5 US gpm).

In addition, there is an EMO located on the front panel of the 19” cabinet supplied with this unit.

All three of these elements are connected in series and are wired to the Eaton PC975-2689 Power Distribution Unit (refer to section 12 for wiring details).

In the event of any one of these interlock conditions being triggered, the power is cut from the Kepco power supplies and the system is de-energized. The Eaton PDU will have to be manually reset before operation. If the EMO is used for shut down, then this will also have to be manually reset. The temperature switch and flow switch automatically reset once the proper operating conditions have been met.

|  | In the event of an interlock tripping it is strongly recommended that the root cause of the trip be determined. In the event that the root cause cannot be determined please contact Tom King (tom@gmw.com) for technical support. |
| --- | --- |

## Cooling

The Model 5203 can be operated up to an average coil temperature of 80°C.  Assuming an ambient laboratory temperature of 20°C and a temperature coefficient of resistivity for copper of 0.0039/°C, the hot resistance of the coil should not exceed 25% more than the ambient temperature "cold" resistance.  The coil thermostat will open when any coil cooling plate temperature exceeds approximately 80°C . Clean, cool (16°C - 20°C) water at 8 l/min and 4 bar (60 psid) must be used to cool the 5203 magnet.

The copper cooling tubes are electrically isolated from the coils to avoid electrochemical corrosion. A 50 micron filter should be placed before the input to the magnet to trap particulates.

Configuration Suggestion: For continuous operation of the magnet it may be appropriate to use a recirculating chiller to reduce water and drainage costs. The chiller capacity will depend on whether cooling is required for the magnet alone or magnet and power supply. For the Model 5203 Electromagnet with two Kepco BOP20-50 power supplies a suitable chiller is the Thermo Fisher TF9 recirculating water chiller (900W).

Do not use corrosion inhibitors in high quality electrical systems since the water conductivity is increased which can result in increased leakage currents and electrochemical corrosion.

Avoid cooling the magnet below the dew point of the ambient air.  Condensation may cause electrical shorts and corrosion.

During operation the resistance can be checked using a voltmeter across the magnet.  The voltage will rise to a constant value once thermal equilibrium has been reached.  If it is desired to save water, the flow can be reduced until the hot resistance is approached.  NOTE: This adjustment must be made slowly enough to allow for the thermal inertia of the coils.

## Mounting Plate

# OPERATION

## General

|  | The 5203 electromagnet is water cooled and should not be operated without water. Attempts to operate the magnet as an air cooled system will most likely result in irreversible damage to the magnet. |
| --- | --- |
|  | This product is an electromagnet and creates significant magnetic fields both within the magnet and in the surrounding area. Make sure that there are no ferromagnetic items, tools or components loose in a 1 meter radius surrounding the magnet. |
|  | The Model 5203 is not a shielded electromagnet. In operation the magnet fringing field can be in excess of 0.5mT (5G). This can cause malfunctioning of heart pacemakers and other medical implants. We recommend that the fringing field should be mapped and warning signs be placed outside the 0.5mT (5G) contour. Entry to this region of higher field should be restricted to qualified personnel |
|  | Do not move magnetically sensitive items into the close vicinity of the magnet. Even some anti-magnetic watches can be damaged when placed in close proximity to the magnet during operation. Credit cards, and magnetic disks are affected by magnetic fields as low as 0.5mT (5G). |

The magnet operates as a conventional water cooled electromagnet.

- Adjust the cooling water flow to approximately 8 liters/min (2.1 US gpm).  For operation at less than maximum power the water flow may be correspondingly reduced.  Note that the inlet water temperature will determine the actual flow rate required.  The above specified flow rates were determined with a water inlet temperature of approximately 18°C.

- Turn on the power supply and increase the current until the desired field is reached. Note that constant current operation is preferred for stable field performance. Please refer to the user manual of the selected power supply.

## Field Control Operation

The necessity to use calibration curves can be avoided by using a field controller to sense the magnetic field and provide a corresponding power supply control signal through the power supply programming inputs.  Contact GMW for suitable instrumentation and software (Labview) (sales@gmw.com).

# MAINTENANCE

|  | Whenever performing maintenance on this magnet all electrical plug socket disconnects must be disconnected and a plug lockout device used to properly shut down the electrical system. Refer to Lockout/Tagout procedure defined in section 2.2 of the “Field Magnet for Drive Testing” User Manual. |
| --- | --- |

## Cooling Circuit

Check the cooling water circuit to ensure the water is clean and free of debris and bacterial growth.  If an in-line water filter is used, ensure the filter is clean. If this magnet is used in conjunction with a closed cycle chiller please refer to the User Manual for details specific to that chiller for cleaning and maintenance.

## Strain Relief Check

During the scheduled maintenance it is recommended that cables at the base of the 5203 Projected Vertical Field Electromagnet be inspected (Refer to section 12.4 cable 16907-0116-1-S1_A). In particular the strain relief should be checked to confirm that the cables are properly restrained from movement. The Lockout/Tagout procedure from section 1.3 should be followed during this procedure.

|  | When inspecting cables to the magnet the magnet should be de-energized to zero amps and the power supply switched off. Refer to drawing 16907-0116-1-S1_A in section 12.4. |
| --- | --- |

## Preventative Maintenance

The interlocks should be cycled periodically, recommendation at least once per quarter.

# STANDARD OPTIONS

## Standard Pole Options

The 5203 electromagnet is supplied with Pole A as standard. Standard Poles B, C and D can be ordered additionally to provide different performance characteristics. See Section 9 for performance characteristics of poles A, B, C and D.

# CUSTOM OPTIONS

The model 5203 electromagnet can have customized performance by means of switching out the standard pole with a customized pole. GMW Associates will work with customers to provide the magnetic design of custom poles and will also manufacture such poles. Please direct inquiries to sales@gmw.com

# EXCITATION CURVES AND FIELD UNIFORMITY

Excitation curves for the standard pole Type A

Near and far vertical field profiles for standard pole Type A.

| z=1mm | Br/Bz Over Disc Ø4mm | Bz Over Disc Ø4mm | Peak Bz on axis [T] |
| --- | --- | --- | --- |
| Pole A | ± 7.54 % | ± 2.36 % | 0.42 |
| Pole B | ± 7.79 % | ± 3.69 % | 0.50 |
| Pole C | ± 6.72 % | ± 3.76 % | 0.61 |
| Pole D | ± 2.38 % | ± 7.35 % | 1.05 |

| z=2mm | Br/Bz Over Disc Ø4mm | Bz Over Disc Ø4mm | Peak Bz on axis [T] |
| --- | --- | --- | --- |
| Pole A | ± 4.12 % | ± 2.29 % | 0.48 |
| Pole B | ± 3.33 % | ± 2.91 % | 0.57 |
| Pole C | ± 2.36 % | ± 3.07 % | 0.67 |
| Pole D | ± 4.87 % | ± 1.67 % | 1.05 |

| z=3mm | Br/Bz Over Disc Ø4mm | Bz Over Disc Ø4mm | Peak Bz on axis [T] |
| --- | --- | --- | --- |
| Pole A | ± 1.68 % | ± 1.82 % | 0.51 |
| Pole B | ± 0.64 % | ± 2.03 % | 0.60 |
| Pole C | ± 0.42 % | ± 2.06 % | 0.69 |
| Pole D | ± 5.81 % | ± 0.17 % | 0.95 |

| z=4mm | Br/Bz Over Disc Ø4mm | Bz Over Disc Ø4mm | Peak Bz on axis [T] |
| --- | --- | --- | --- |
| Pole A | ± 0.13 % | ± 1.28 % | 0.52 |
| Pole B | ± 1.00 % | ± 1.24 % | 0.60 |
| Pole C | ± 2.07 % | ± 1.17 % | 0.68 |
| Pole D | ± 6.07 % | ± 0.65 % | 0.85 |

| z=5mm | Br/Bz Over Disc Ø4mm | Bz Over Disc Ø4mm | Peak Bz on axis [T] |
| --- | --- | --- | --- |
| Pole A | ± 1.00 % | ± 0.83 % | 0.52 |
| Pole B | ± 2.00 % | ± 0.70 % | 0.58 |
| Pole C | ± 3.00 % | ± 0.51 % | 0.65 |
| Pole D | ± 5.95 % | ± 0.89 % | 0.75 |

| z=6mm | Br/Bz Over Disc Ø4mm | Bz Over Disc Ø4mm | Peak Bz on axis [T] |
| --- | --- | --- | --- |
| Pole A | ± 1.74 % | ± 0.55 % | 0.50 |
| Pole B | ± 2.64 % | ± 0.40 % | 0.56 |
| Pole C | ± 3.57 % | ± 0.21 % | 0.60 |
| Pole D | ± 5.71 % | ± 0.90 % | 0.67 |

| z=7mm | Br/Bz Over Disc Ø4mm | Bz Over Disc Ø4mm | Peak Bz on axis [T] |
| --- | --- | --- | --- |
| Pole A | ± 2.28 % | ± 0.34 % | 0.48 |
| Pole B | ± 3.08 % | ± 0.16 % | 0.53 |
| Pole C | ± 3.91 % | ± 0.08 % | 0.56 |
| Pole D | ± 5.46 % | ± 0.83 % | 0.59 |

| z=8mm | Br/Bz Over Disc Ø4mm | Bz Over Disc Ø4mm | Peak Bz on axis [T] |
| --- | --- | --- | --- |
| Pole A | ± 2.63 % | ± 0.18 % | 0.46 |
| Pole B | ± 3.28 % | ± 0.27 % | 0.49 |
| Pole C | ± 4.03 % | ± 0.44 % | 0.52 |
| Pole D | ± 5.16 % | ± 1.01 % | 0.53 |

| z=9mm | Br/Bz Over Disc Ø4mm | Bz Over Disc Ø4mm | Peak Bz on axis [T] |
| --- | --- | --- | --- |
| Pole A | ± 2.89 % | ± 0.15 % | 0.43 |
| Pole B | ± 3.50 % | ± 0.20 % | 0.46 |
| Pole C | ± 4.14 % | ± 0.41 % | 0.47 |
| Pole D | ± 4.94 % | ± 0.87 % | 0.47 |

| z=10mm | Br/Bz Over Disc Ø4mm | Bz Over Disc Ø4mm | Peak Bz on axis [T] |
| --- | --- | --- | --- |
| Pole A | ± 3.14 % | ± 0.18 % | 0.41 |
| Pole B | ± 3.61 % | ± 0.28 % | 0.42 |
| Pole C | ± 4.11 % | ± 0.41 % | 0.43 |
| Pole D | ± 4.77 % | ± 0.70 % | 0.43 |

# Block Diagram

# WIRING DRAWINGS

## Drawing 13907-0025-1_B 5203 1xBOP20-50 and PDU System Wiring

## Drawing 13907-0026-1_B 5203 2xBOP20-50 and PDU System Wiring

## Drawing 16907-0116-1-S1_A Cable Current Interlock 5203 and PDU to 
BOP20-50

# APPENDIX

## Recommended Mains Connection

GMW recommends the following Mains Connection as per conformance with NFPA 79

- MELTRIC DS30 Receptacle (female):	33-34167

- MELTRIC DS30 Inlet (male):		33-38167

- Additional MELTRIC DS30 accessories as needed.